## Slide 1: Get Acquainted with the Latest Versions of LC & SMBA
USGS, TRDI, and SonTek Versions
David S. Mueller, OSW
April, 2013

## Slide 2: Outline
Collecting Data in Moving-Bed Conditions
Versions available for LC and SMBA
Brief review of LC basics
New LC quality assurance checks
Brief review of SMBA basics
Using TRDI versions of LC and SMBA
Using SonTek version of LC and SMBA
Summary

## Slide 3: Collecting Data in Moving-Bed Conditions
Use DGPS if possible
  LC & SMBA are NOT substitutes for using GPS
  Webinar recorded last month on use of GPS
When to use LC Correction
  GPS use is not valid or possible
  BT and accurate compass heading can be maintained
When to use SMBA Correction
  GPS use is not valid or possible
  Accurate compass headings cannot be maintained through cross section
  BT is invalid intermittently
When to use Mid-Section technique
  Criteria for GPS, Loop Correction, or Stationary Correction not met

## Slide 4: Available Versions of LC & SMBA
USGS
SonTek RiverSurveyor Live
TRDI WinRiver 2

## Slide 5: Review of Loop Basics
Compass calibration and accuracy throughout the cross section is critical
Bottom track must be maintained
Uniform boat speed
Don’t loiter at edges
Return to the exact starting position
Loop duration should be 3 minutes or greater (longer=better)
May not be accurate for water velocities less than 0.8 ft/s
Flow
Distance moved upstream
DUP

## Slide 6: Previous Criteria in LC
Moving-bed velocity > 0.04 ft/s
Closure error must be in upstream direction
  Water direction +/- 135 degrees
Moving-bed velocity / Mean water velocity > 0.01
Warnings provided for:
  Percent bad bottom track
  Potential compass error
Determination of loop validity left to the user

## Slide 7: New LC Quality Assurance Checks
Navigation must be set to Bottom Track
  Message: “DATA FILE NOT REFERENCED TO BOTTOM TRACK. Reprocess file with reference set to bottom track.”
RSL: Composite tracks must be turned off
  Message: “Composite tracks are on. The composite tracks feature in RiverSurveyor should be disabled and the Matlab file recreated in RiverSurveyor and then reloaded in LC.”
Slow water velocity warning
  Message: “WARNING: The water velocity is less than 0.8 ft/s and could cause the loop method to be inaccurate. CONSIDER USING A STATIONARY TEST TO CHECK MOVING-BED CONDITIONS.”

## Slide 8: Invalid Bottom Track
Invalid bottom track > 5%
  Message: “WARNING: Percentage of bad bottom track values exceeds 5. Loop may not be accurate. PLEASE REVIEW DATA.”
Invalid bottom track > 20%
  Message: “ERROR: Percent invalid Bottom Track exceeds 20 percent. THE LOOP IS NOT ACCURATE. TRY A STATIONARY MOVING-BED TEST.”
Consecutive invalid bottom track > 9 seconds
  Message: “ERROR: Consecutive invalid Bottom Track exceeds 9 seconds. THE LOOP IS NOT ACCURATE. TRY A STATIONARY MOVING-BED TEST.”

## Slide 9: Compass Evaluation
Find apex of loop
Compute mean and standard deviation of flow direction for each leg
  Difference reported to user
Compute uncertainty associated with difference in flow direction
  95% uncertainty ~ stdDev1 * 2 + stdDev2 * 2
Compute potential error caused by compass error
Compass ERROR reported
  Difference flow directions > 3 degrees
  Difference in flow directions > 95% uncertainty
  Potential error > 5%
Width
Closure error
due to compass
Difference in flow direction
Vmb due to compass

## Slide 10: Compass Error Message
ERROR: Difference in flow direction between out and back sections of
loop could result in a 5 percent or greater error in final discharge.
REPEAT LOOP AFTER COMPASS CAL. OR USE A STATIONARY MOVING-BED TEST.

## Slide 11: Closure & General Errors
Downstream closure & moving-bed velocity > 0.04 ft/s
  Message: “ERROR: Loop Closure Error not in Upstream Direction -- REPEAT LOOP or USE STATIONARY TEST.”
  No error message if moving-bed velocity <= 0.04 ft/s
Errors have been identified
  Message: “ERROR: Due to ERRORS noted above this loop is NOT VALID. PLEASE CONSIDER SUGGESTIONS.”

## Slide 12: Review of SMBA
Initially developed for StreamPro
  False moving-bed when swimming
Applicable to all ADCPs
Uses mean velocity direction to determine upstream movement on an ensemble by ensemble basis

## Slide 13: Correction Methods
LC: The moving-bed velocity from the loop moving-bed test is distributed to each ensemble based on the ratio of the near-bed water speed in each ensemble to the mean of near-bed water speeds for the whole transect.
  WinRiver II calls this the distributed method
SMBA: Computes a linear coefficient (C) relating the near-bed water speed to moving-bed speed from the stationary tests (Vmb=C*Vnb). This equation is then used to distribute the moving-bed correction to each ensemble in the transect.
  WinRiver II calls this the proportional method

## Slide 14: Collecting MBT in WinRiver II

## Slide 15: Add a File as Moving-Bed Test
Right Click on the mmt filename
Select desired type of test
From file dialog select file to add
This must be done for MBTs collected using the StreamPro with a PDA

## Slide 16: Processing MBT
MBTs are located under the QA/QC – Moving Bed Test node
Reprocess / Evaluate
View Results
View Summary

## Slide 17: Reprocess to Manually Evaluate Data
Loop
Stationary

## Slide 18: FAIL / PASS – Still valid Loop MBT

## Slide 19: Test Result with Error

## Slide 20: Select Tests
Check to use
Uncheck to not use
If test is invalid (ERROR), WR2 will not use the results even if checked
Only loops that meet the criteria for a moving-bed condition will be applied.
All positive stationary test will be applied even is less than 1% of mean velocity.

## Slide 21: MBT Summary with Error

## Slide 22: MBT Summary with valid MBT
Multiple loops can be used.

## Slide 23: Stationary MBT Result
If Stationary is checked it is applied even if less than 1% error.

## Slide 24: Example of Stationary MBT < 1%

## Slide 25: Multiple Tests

## Slide 26: Correction Type
Distributed
  Same as USGS LC
  With multiple Loop MBTs average of results is used
Proportional
  Similar to USGS SMBA
  No negative corrections are applied
  Correction is always applied even if less than 1%
  Used with Stationary or mix of Stationary and Loop

## Slide 27: Where are Corrections Applied
Discharge corrections based on the results of moving bed tests are only applicable when Bottom Track is the selected navigation reference.
The corrected discharge is displayed in the MBT Summary and in the Composite Tabular display, the Discharge Summary, the Total Discharge Time Series, Q Measurement Summary, and in HydroML output.
Moving bed corrections are NOT applied to any other parameters or outputs such as velocity and discharge contour, profile, and time series plots, or the Classic or Generic ASCII outputs.

## Slide 28: BT with Loop Checked

## Slide 29: GPS with Loop Checked (Bug?)
Discharge is referenced to GPS correctly!!

## Slide 30: Differences with USGS LC & SMBA
Rounding in ASCII output file may result in small differences in resulting computed discharge (< 1%)
WinRiver II behavior not in USGS codes:
  Multiple loops can be averaged
  Loops and stationary tests can be applied together
  Stationary test is always applied even if correction is < 1%
  Negative stationary tests are ignored in WR2

## Slide 31: Loop MBT with RiverSurveyors
Due to potential compass calibration and accuracy issues associated with the RiverSurveyor M9/S5 ADCPs, use of the stationary moving-bed test will be consistently more accurate than the loop test and is recommended as the preferred moving-bed test for RiverSurveyors until the compass calibration issues can be completely resolved (http://hydroacoustics.usgs.gov/movingboat/pdfs/Best_Practice_for_Calibrating_SonTek_RiverSurveyor.pdf)
The OSW is working with SonTek to try and resolve the compass accuracy issue.
This webinar will show how to use RiverSurveyor Live with the loop MBT so that it can be properly applied when the compass issue is resolve.

## Slide 32: Collecting MBT in RiverSurveyor Live

## Slide 33: System Tab

## Slide 34: Loop MBT Collection Dialog

## Slide 35: Stationary MBT Collection Dialog

## Slide 36: Loading MBT
First open your transect files
Then open your MBT
Only 1 loop can be open at a time
Multiple stationary tests can be opened

## Slide 37: Dialog for Loop MBT

## Slide 38: Dialog for Stationary MBT

## Slide 39: Applying MBT for Q Correction
Stationary MBT Correction
Loop MBT Correction

## Slide 40: Results of Apply MBT Correction
If these columns are not present right click on heading and select them from the list

## Slide 41: SMBA Correction not Displayed (bug?)
Samples Tab

## Slide 42: Composite Tracks
Composite Tracks should be turned off when using LC or SMBA
Bug in RSL
  Uncheck composite tracks in Processing Toolbox – Thresholds and click Apply All
  Return to Processing Toolbox – System Settings and with LC or SM selected click Apply All

## Slide 43: RSL’s LC & SMBA
Follows USGS code very closely
Corrections computed using USGS LC & SMBA formulas
May be some differences in final Q due to rounding differences
Potential issue:
  Invalid BT is identified by repeated identical BT East Velocity
  Recent data showed valid BT with identical BT East Velocity
  USGS LC 4.10 (Beta) has fixed this problem
  Issue will be reported to SonTek
  Only causes error in reported % bad BT, actual results of loop are unaffected

## Slide 44: Latest Versions
USGS
  LC 4.04 available to public
  LC 4.10 being test by HaWG
    Speed improvements
    Fixed invalid BT identification for RSL data
  SMBA 6.13
    Speed improvements
TRDI
  WinRiver II 2.10
SonTek
  RiverSurveyor Live 3.60

## Slide 45: REMEMBER: LC & SMBA are not substitutes for using GPS

## Slide 46: Summary
Use GPS, if possible, for moving-bed conditions
Use Stationary MBT when compass is not accurate throughout the cross section
Evaluate and set extrapolation prior to creating files for use by USGS LC or SMBA
Use any of the implementations of LC & SMBA (WinRiver II, RiverSurveyor Live, or USGS)

## Slide 47: COMMENTS & QUESTIONS