## Slide 1: Review and Rating Discharge Measurements
David S. Mueller
Office of Surface Water
March 2010

## Slide 2: Overview
Considerations in Rating a Measurement
Six Review Steps
  Gather information for rating as you review
Three Examples

## Slide 3: Rating Process
Quantitative assessment of measurement uncertainty
  Currently being worked on
Current practice is subjective
  Where we can quantify we will
  Where we can’t we will make qualitative estimates

## Slide 4: Rating Considerations
Variability of transect discharges
Pattern in transect discharges
Accuracy of top and bottom extrapolation
Accuracy of edge extrapolation
Quantity and distribution of missing data
Quality of boat navigation reference
  Bias or noise in bottom track
  Noise in GPS data

## Slide 5: Spread in Transect Discharges
A coefficient of variation (COV) is computed by the software (Std./|Avg.|)
The COV represents the sampled random error in the discharge measurement.
Dividing the COV by the sqrt of the number of passes and multiplying by the correct value will provide the 95% uncertainty associated with random errors as measured
  For 4 passes: 1.6*COV
  For 8 passes: 0.8*COV
  This does not include any bias errors
  This is probably and optimistic estimate

## Slide 6: Example of Random Error
95% Uncertainy (Random Error) = 0.05*0.8 = 4%

## Slide 7: Patterns in Transects
Random errors should not display a pattern
A pattern in flow may indicate the need to adjust measurement procedures
  Rapidly rising or falling – average fewer transects
  Periodic – use more transects to capture variability

## Slide 8: Top Extrapolation
Average appropriate number of ensembles
Evaluate discharge profile
Check field notes for wind effects
Try different extrapolation methods and compare change in total discharge
  Little change – accurate extrapolation likely
  Significant change – how good is the data on which the selection is made
    Surface effects captured?
    Field notes consistent with pattern?
What percentage of flow is in the top portion
Subjective judgment

## Slide 9: Bottom Extrapolation
Smooth streambed
  Accurate extrapolation likely
Rough streambed
  Extrapolation uncertain

## Slide 10: Edge Estimates
What percentage of flow is in the edges?
  May be small enough to disregard accuracy
  Recommend no more than 5% in each edge
Subjective judgment
  Large edge estimates
  Irregular edge shapes
Edge Q = 25% of total

## Slide 11: Invalid Data
Percent of invalid data
Distribution of invalid data
How well is that portion of the flow estimated

## Slide 12: Summary of Rating Approach
Consider the data
Consider the hydraulics
Consider site conditions
Use statistics, if possible
Use your judgment
Pay attention to potential errors (as a percentage) as you review the data

## Slide 13: Steps of Review
QA/QC
Composite tabular
Stick Ship Track
Velocity contour
Extrapolation
Discharge summary

## Slide 14: QA/QC
ADCP Test Error Free?
Compass Calibration
  Calibration and evaluation completed
  < 1 degree error
  Important for Loop and GPS
Moving-Bed Test
  Loop
    Requires compass calibration
    Use LC to evaluate
  Stationary
    Duration and Deployment
    SMBA required to evaluate StreamPro

## Slide 15: Composite Tabular
Number of Ensembles
  > 150 ???
Lost Ensembles
  Communications error
Bad Ensembles
  Distribution more important than percentage
  Distribution evaluated in velocity contour plot
Reasonable temperature

## Slide 16: Stick Ship Track
Check with appropriate reference (BT, GGA, VTG)
Track and sticks in correct direction
  Magnetic variation
  Heading time series
Sticks representative of flow
  WT thresholds
Irregular ship track?
  BT thresholds
  BT 4-beam solutions

## Slide 17: Velocity Contour
Patterns in velocity distribution
  Nonuniform
  Unnatural
  WT Thresholds
Spikes in bottom profile
  Turn on Screen Depth
Distribution of invalid and unmeasured areas
  Qualitative assessment of uncertainty

## Slide 18: Velocity Extrapolation
Average ensembles (10-20)
Evaluate several transects
Use Power/Power as default
  Find reasons to deviate from Power/Power
  Don’t waste time where there is little difference
    In doubt try different options an evaluate change in discharge
    Helps evaluate potential uncertainty from extrapolation
  Person collecting data gets benefit of doubt

## Slide 19: Discharge Summary
Transects within 5%
Directional bias
  GPS used
    Magvar correction
Consistency
  Discharges (top, meas., bottom, left, right)
  Width, Area
  Boat speed, Flow Speed, Flow Dir.

## Slide 20: Example 1 – Compass Cal.
Compass Evaluation is Good!

## Slide 21: Example 1 – Moving-Bed Test
Loop looks good
LC finds no problems with loop
  Lost BT
  Compass Error
  No moving-bed correction required

## Slide 22: Example 1 – Composite Tabular
> 300 ensembles
No lost ensembles
0-1 bad ensembles
Temperature reasonable

## Slide 23: Example 1 – Ship Track / Contour
Ship track consistent
No spikes in streambed in contour plot
No hydraulically unusual patterns
Unmeasured areas reasonable in size

## Slide 24: Example 1 - Extrapolation
Top extrapolation a bit uncertain
Try Power/Power
  Changed Q = 0.6%
Const / No Slip = OK

## Slide 25: Example 1 – Discharge Summary
All within 5%
1 negative Q in left edge
Everything else very consistent

## Slide 26: Example 1 - Rating
Base uncertainty
  1.6*0.01= 1.6%
Invalid Data
  None
Extrapolation
  0.5-1% uncertainty in extrapolation
  Top is only about 10% of total
Edges
  About 1% of total Q
Final assessment
  Rate this measurement “GOOD”

## Slide 27: Example 2 – QA/QC
No test errors
Compass Cal
  High pitch and roll
Moving-bed tests
  Bottom track problems

## Slide 28: Example 2 – Bottom Track Issue
Time series plots show spikes
  3-beam solution
  High vertical velocity

## Slide 29: Example 2 – Composite Tabular
Too few ensembles
Short duration
Bad ensembles

## Slide 30: Example 2 – Ship Track / Contour
Irregular ship track
  Thresholds didn’t help
No spikes in streambed
Invalid data
  Not many
  High percentage because of so few ensembles
  Estimates probably reasonable
Large unmeasured areas
  % measured < 20%

## Slide 31: Example 2 - Extrapolation
Averaged 10
Highly variable
Difference between Power/Power and Constant / No Slip is 3.6%
  Uncertain about extrapolation accuracy

## Slide 32: Example 2 – Discharge Summary
Not within 5%
  8 passes
Consistent edges
Inconsistent width and area
  Bottom track problems
High Q COV

## Slide 33: Example 2 - Rating
Base uncertainty
  0.14 * 0.8 = 10%
Invalid data
  Some but not significant
  Too few ensembles (or too few transects)
  Bottom track problems
  Large unmeasured areas
Extrapolation
  Power/Power to Constant/No Slip = 3.6%
Edges
  Reasonable
Final Assessment
  Rate this measurement “POOR”

## Slide 34: Example 3 – QA/QC
No errors in ADCP test
Good compass calibration
Two loop tests

## Slide 35: Example 3 – Composite Tabular
High percentage of bad ensembles

## Slide 36: Example 3 – Composite Tabular
Very few bad ensembles with reference set to GGA!!

## Slide 37: Example 3 – Ship Track / Contour
Everything looks good with reference set to GGA

## Slide 38: Example 3 – Ship Track / Contour
With bottom track we need to qualitatively assess the effect of the invalid data on uncertainty

## Slide 39: Example 3 - Extrapolation
Given the variability among the profile plots constant / no slip seems reasonable.
Changing to power / power made no difference in the final Q
Might justify power / power

## Slide 40: Example 3 – Discharge Summary
Small directional bias
  Typical of magvar or compass error with GPS
  Adjusting magvar 2.5 degrees removed directional bias and changed Q by 1.3 cfs.
Edge discharge are inconsistent but small relative to total Q

## Slide 41: Example 3 – Discharge Summary
Bottom track referenced discharge are lower indicating a moving bed was present.
Edge discharges are inconsistent but small
Would require correction with LC

## Slide 42: Example 3 - Rating
GPS (GGA or VTG) should be used as the reference.
Base uncertainty
  0.03 * 1.6 = 4.8%
  Value inflated due to directional bias
No significant invalid data
Extrapolation and edges would make little difference
Final assessment
  Could rate this measurement “Good” wouldn’t argue much if rated “Fair”

## Slide 43: Example 3 - Rating
If GPS were not available and BT had to be used the rating would change.
Loops
  High percentage of invalid bottom track
  5-8% moving-bed bias
Base uncertainty
  0.02 * 1.4 = 3.2%
Invalid data
  Represent a significant percentage of the flow
  Likely estimated well from neighboring data
Edges
  Inconsistent
  Small percentage
Extrapolation
  Errors very small
Final assessment
  Loop correction probably within 4% + 3.2% random error
  I would probably rate this measurement “Fair”

## Slide 44: Summary
Consider potential uncertainty or errors as you review the data
Keep review to a minimum unless errors are identified or suspected
Consider the percent of effect of an error or uncertainty will have on the discharge
Extrapolation uncertainty can be large in shallow streams
Use the COV in the discharge summary as the base minimum uncertainty
Contact OSW if you have unusual data or you have questions

## Slide 45: Questions???